## GLP-1 Validation
GLP-1: Validation (Ansh New version: Dec-2019 kit)
3-3-2020: Radha
Precision: | [6 empty] | Specificity: |  |  |  | Dilution-Linearity: | [9 empty] | Spike-Recovery:
Plate-1 | [18 empty] | Assay background:
|  | GLP-1 (pg/ml) | Mean | St Dev | SEM | CV |  |  | GLP-1 | % Cross-reactivity |  |  | GLP-1 |  |  |  | GLP-1 |  |  | GLP-1 |  |  | GLP-1
Pool-1 | x, 50 ul | 84.923 | [5 empty] | Concentration | pg/ml |  |  |  | pg/ml | % Recovery |  |  | pg/ml | % Recovery |  | pg/ml |  |  | pg/ml | % Recovery
| x, 50 ul | 80.562 |  |  |  |  | Normal assay |  | 73.49 |  | Pool-1 | X | 55.245 |  | Pool-A | X | 75.48 |  | Extraction buffer * | < 1.90 | Pool-A | X/2 Sample | 18.95
| x, 50 ul | 81.69 |  |  |  |  | MPGF-2 | 0.194 ng/ml | 70.06 | Non-detectable |  | X/2 dil. | 22.091 | 80% |  | X/2 dil. | 33.85 | 90% | PBS * | < 1.90 |  | X/2 QC2 | 49.243
| x, 50 ul | 79.466 | 81.6603 | 2.3571 | 1.0523 | 2.8864 |  | 0.39 ng/ml | 73.01 | Non-detectable |  | X/4 dil | 12.037 | 109% |  | X/4 dil | 17.54 | 103% | Standard-A * | < 1.90 |  | X/2 Sample + X/2 QC2 | 62.639 | 91.8%
Pool-A | x, 50 ul | 75.392 | [5 empty] | 0.781 ng/ml | 70.98 | Non-detectable |  | X/8 dil. | 3.985 | 66.3% |  | X/8 dil. | 7.82 | 89% | * Processed as samples |  | Pool-B | X/2 Sample | 38.209
| x, 50 ul | 77.04 | [5 empty] | 1.562 ng/ml | 71.48 | Non-detectable |  | X/16 dil. | < 1.90 |  |  | X/16 dil. | 3.16 | 81% |  |  |  | X/2 QC2 | 50.448
| x, 50 ul | 73.108 | [5 empty] | 3.125 ng/ml | 70.12 | Non-detectable |  | X/32 dil. | < 1.90 |  |  | X/32 dil. | < 2.00 |  |  |  |  | X/2 Sample + X/2 QC2 | 81.022 | 91.4%
| x, 50 ul | 76.391 | 75.4828 | 1.7222 | 0.76883 | 2.2815 |  | 6.25 ng/ml | 74.01 | Non-detectable |  | X/64 dil. | < 1.90 |  | Pool-B | X | 67.55 |  |  |  | Pool-C | X/2 Sample | 28.458
Pool-B | x, 50 ul | 67.036 | [5 empty] | 12.5 ng/ml | 79.36 | 8% |  | X/128 dil. | < 1.90 |  |  | X/2 dil. | 30.17 | 89% |  |  |  | X/2 QC2 | 48.989
| x, 50 ul | 67.986 | [5 empty] | 25 ng/ml | 98.36 | 34% | Pool-2 | X | 107.871 |  |  | X/4 dil | 14.52 | 96% |  |  |  | X/2 Sample + X/2 QC2 | 72.589 | 93.7%
| x, 50 ul | 68.159 | [5 empty] | 50 ng/ml | 190.16 | 159% |  | X/2 dil. | 48.251 | 89.4% |  | X/8 dil. | 5.99 | 83%
| x, 50 ul | 67.008 | 67.5472 | 0.61071 | 0.27264 | 0.90413 | OXM | 10 ng/ml | 71.61 | Non-detectable |  | X/4 dil | 22.118 | 91.6% |  | X/16 dil. | < 2.00
Plate-2 | [6 empty] | Glucagon | 10 ng/ml | 72.45 | Non-detectable |  | X/8 dil. | 11.046 | 99.9% |  | X/32 dil. | < 2.00
|  | GLP-1 (pg/ml) | Mean | St Dev | SEM | CV | Glicentin | 10 ng/ml | 70.62 | Non-detectable |  | X/16 dil. | 3.499 | 63.3% | Pool-C | X | 61.9
Pool-C | x, 50 ul | 63.053 |  |  |  |  | GLP-2 | 10 ng/ml | 71.34 | Non-detectable |  | X/32 dil. | < 1.90 |  |  | X/2 dil. | 26.98 | 87% | [5 empty] | QC-1: 45.55pg/ml (Exptd: 45 +/- 11 pg/ml)
| x, 50 ul | 61.301 |  |  |  |  | Insulin | 10 ng/ml | 72.9 | Non-detectable |  | X/64 dil. | < 1.90 |  |  | X/4 dil | 12.5 | 93% | [5 empty] | QC-2: 118.93pg/ml (Exptd: 120 +/- 24pg/ml)
| x, 50 ul | 62.056 | [9 empty] | X/128 dil. | < 1.90 |  |  | X/8 dil. | 4.71 | 75%
| x, 50 ul | 61.204 | 61.9035 | 0.85575 | 0.38203 | 1.3824 |  |  |  |  | Pool-3 | X | 108.692 |  |  | X/16 dil. | < 2.00 | [6 empty] | Stds pg/mL | BackCal. | OD | AvgOD | CV
Pool-D | x, 50 ul | 59.952 | [9 empty] | X/2 dil. | 45.8 | 84.2% |  | X/32 dil. | < 2.00 | [6 empty] | 0 | Range? | 0.059 | 0.059 | 1.7
| x, 50 ul | 57.805 | [9 empty] | X/4 dil | 22.104 | 96.5% | Pool-D | X | 59.12 | [7 empty] | Range? | 0.058
| x, 50 ul | 59.383 | [9 empty] | X/8 dil. | 10.95 | 99.1% |  | X/2 dil. | 26.05 | 88% | [5 empty] | 15 | 15.804 | 0.093 | 0.092 | 0.9
| x, 50 ul | 59.315 | 59.1138 | 0.91806 | 0.40985 | 1.553 | [5 empty] | X/16 dil. | 3.749 | 68.5% |  | X/4 dil | 11.68 | 90% | [6 empty] | 15.466 | 0.091
Plate-3 | [11 empty] | X/32 dil. | < 1.90 |  |  | X/8 dil. | 3.24 | 56% | [5 empty] | 38 | 38.293 | 0.224 | 0.227 | 2.1
|  | GLP-1 (pg/ml) | Mean | St Dev | SEM | CV | [5 empty] | X/64 dil. | < 1.90 |  |  | X/16 dil. | < 2.00 | [7 empty] | 39.134 | 0.231
Pool-E | x, 50 ul | 97.04 | [9 empty] | X/128 dil. | < 1.90 |  |  | X/32 dil. | < 2.00 | [6 empty] | 84 | 83.006 | 0.749 | 0.758 | 1.7
| x, 50 ul | 102.124 | [8 empty] | Pool-4 | X | 56.644 |  | Pool-E | X | 99.58 | [7 empty] | 84.209 | 0.767
| x, 50 ul | 100.217 | [9 empty] | X/2 dil. | 25.041 | 88.4% |  | X/2 dil. | 43.53 | 87% | [5 empty] | 125 | 124.244 | 1.464 | 1.481 | 1.7
| x, 50 ul | 98.969 | 99.5875 | 2.1372 | 0.95409 | 2.146 | [5 empty] | X/4 dil | 13.108 | 104.7% |  | X/4 dil | 20.92 | 96% | [6 empty] | 126.053 | 1.499
Pool-F | x, 50 ul | 56.234 | [9 empty] | X/8 dil. | 3.927 | 60.1% |  | X/8 dil. | 9.98 | 95% | [5 empty] | 215 | 214.328 | 3.357 | 3.371 | 0.6
| x, 50 ul | 58.159 | [9 empty] | X/16 dil. | 1.768 | 90.1% |  | X/16 dil. | 2.11 | 42% | [6 empty] | 215.646 | 3.385
| x, 50 ul | 58.398 | [9 empty] | X/32 dil. | < 1.90 |  |  | X/32 dil. | < 2.00
| x, 50 ul | 57.692 | 57.6208 | 0.96988 | 0.43298 | 1.6832 | [5 empty] | X/64 dil. | < 1.90 |  | Pool-F | X | 57.62
[12 empty] | X/128 dil. | < 1.90 |  |  | X/2 dil. | 26.15 | 91%
[16 empty] | X/4 dil | 11.41 | 87%
[16 empty] | X/8 dil. | 2.15 | 38%
[16 empty] | X/16 dil. | < 2.00
[16 empty] | X/32 dil. | < 2.00